## Slide 1
Karl Reich, Ph.D.
Independent Forensics
Lombard IL
A New and Better Forensic Test for
Human Saliva

## Slide 2
Introduction to RSIDTM
Tests for Forensic Laboratories
Lateral Flow Immunochromatographic
Strip Tests for the
Identification of Body Fluids
RSIDTM - SALIVA

## Slide 3
Test Format: Lateral flow immunochromatographic strip test
Antigen Detected: Human Salivary -amylase
Antibodies: 2 Mouse monoclonals, non-overlapping epitopes
Labeling: Colloidal Gold
SALIVA
Sensitivity: < 1 l saliva [No Hook effect]
Cross-reaction: human breast milk, fecal samples (both contain -amylase)
Test Read-out: Visual
Assay Time: 10 minutes ‘active’ time

## Slide 4
SALIVA
15 min
adsorb stain onto swab
extract in provided
extraction buffer
RSIDTM
Running Buffer
mix extract aliquot
with running buffer
(100 l total)
add to RSID-Saliva

## Slide 5
SALIVA
50 nl
250 nl
0.5 l
1.0 l
Saliva
0 l
B/Se/U
Body Fluids
B/Se/Sa/U
0 l
1 l
5 l
10 l
20 l
Extract
Saliva
50 l of
saliva
extracted
in 1 ml

## Slide 6
SALIVA

## Slide 7
SALIVA
Efficient: Assay procedure integrated in DNA-STR analysis
Fast: Test results in 10 minutest post extraction
Sensitive: Detects as little as 50 nl of saliva – stated detection limit: 1 l
No observed cross reaction with blood, urine, sweat, semen, domestic animals,
exotic species,
body fluid mixtures
Minor, but consistent signals seen with breast milk and fecal samples
First specific, non-enzymatic test for salivary amylase
No High Dose Hook effect observed: little or no dilution required,
less possibility of false negatives

## Slide 8
Questions
Comments
Samples
Gripes
Karl Reich
Tel 708.234.1200
Fax 708.978.5115
karl@IFI-test.com
www.IFI-test.com